# Service category 1.1.2

## Basic information

- Quantity				Time scale difference

- Instrument or Artifact			Local clock vs. UTC

- Instrument Type or Method		Comparison against predicted UTC

- Measurand Level or Range		-0.5 s to 0.5 s

- Measurement Conditions

  - Sampling number			100

  - Input Slew Rate			0.5 V/ns

## Instruments

- RF Distribution amplifier		HPDA-15RM-B (SpectraDynamics, Inc.)

- Pulse Distribution amplifier		10188 (TimeTech GmbH)

- Cable					FSJ1-50A (Andrew)

- Time Interval Counter		SR620 (Stanford Research Systems)

## Measurement configuration

## List of uncertainties

- UTC(NMI) Prediction Uncertainty (30days)

- UTC(NMI) Link Uncertainty

- Delay Measurement Uncertainty

- Distribution amplifier Uncertainty

- Cable Uncertainty

- Counter Uncertainty

- Counter channel differential delay

## Details of Uncertainty calculations

- UTC(NMI) Prediction Uncertainty (30days)

  - calculation 1

  - Average			7.0 ns

  - Standard deviation		10 ns

    - Comments

      - Calculated from an average and standard deviation of UTC-UTC(NMI) published in circular T 2010.

      - In this calculation, we assume offset of UTC(NMI) is always zero and fractional deviation is treated as an uncertainty.

  - calculation 2

  - Maximum difference between UTC-UTC(NMI)		23.4 ns

    - Comments

      - We assume the distribution of the maximum value of time difference between UTC-UTC(NMI) as rectangular.

  - calculation 3

  - Standard deviation of the frequency between UTC-UTC(NMI) at 30 days.	2.2E-15

    - Comments

      - When you use this calculation, you have to include last data of UTC-UTC(NMI) in the time scale difference calculation.

- UTC(NMI) Link Uncertainty

  - uA		0.3 ns  (worst value in Circular T published last year)

  - uB		4.8 ns  (worst value in Circular T published last year)

- Delay Measurement Uncertainty

  - 5071A accuracy		1.0E-12

  - Measurement period	1000 s

  - Comments

    - This uncertainty is caused by delay measurement between UTC(NMI) and calibration system using portable clock method.

- Distribution amplifier Uncertainty

10188 (TimeTech GmbH)

  - Temperature coefficient	10 ps/degree

  - Output jitter		20ps

  - Temperature variation	+/-1.0 degree

- Cable Uncertainty

FSJ1-50A (Andrew)

  - Temperature coefficient	0.036 ps/m/degree

  - Temperature variation	+/-5.0 degree

  - Cable length		60 m

- Counter Uncertainty

SR620 (Stanford Research Systems)

  - sampling number		100

  - Timebase error		5E-13 @1s

  - short term stability	5E-13 @1s

  - Input Slew Rate		0.5 V/ns

  - Trigger setting value	1.5 V

  - Time interval		1 s

  - Comments

    - You have to refer manual of your own time interval counter.

- Counter channel differential delay

SR620 (Stanford Research Systems)

  - channel differential delay		500 ps

  - Comments

    - This uncertainty cannot be reduced by averaging.

## Uncertainty Budget table

|  | Source of Uncertainty | Uncertainty ns | Type | Distribution | Conversion factor | Standard Uncertainty ns |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | UTC(NMI) Prediction Uncertainty (30days) | 14 | B | Normal | 1 | 14 |
| 2 | UTC(NMI) Link Uncertainty | 4.8 | B | Normal | 1 | 4.8 |
| 3 | Delay Measurement Uncertainty | 1.00 | A | Normal | 1 | 1.00 |
| 4 | Distribution amplifier Uncertainty | 0.023 | B | Rectangular |  | 0.013 |
| 5 | Cable Uncertainty | 0.02 | B | Rectangular |  | 0.01 |
| 6 | Counter Uncertainty | 0.1 | B | Rectangular |  | 0.06 |
| 7 | Counter channel differential delay | 0.5 | B | Rectangular |  | 0.3 |
| Combined Uncertainty |  |
| Expanded Uncertainty (k=2) | 30 |

## History

- 2011-12-xx first version.