# Service category 1.1.2

## Basic information

- Quantity				Time scale difference

- Instrument or Artifact			Local clock vs. UTC

- Instrument Type or Method		Comparison against post processed UTC

- Measurand Level or Range		-0.5 s to 0.5 s

- Measurement Conditions

  - Sampling number			100

  - Input Slew Rate			0.5 V/ns

## Instruments

  - RF Distribution amplifier		HPDA-15RM-B (SpectraDynamics, Inc.)

  - Pulse Distribution amplifier	10188 (TimeTech GmbH)

  - cable				FSJ1-50A (Andrew)

  - Time Interval Counter		SR620 (Stanford Research Systems)

## Measurement configuration

## List of uncertainties

- UTC(NMI) Link Uncertainty

- Interpolation Uncertainty

- Delay Measurement Uncertainty

- Distribution amplifier Uncertainty

- Cable Uncertainty

- Counter Uncertainty

- Counter channel differential delay

## Details of Uncertainty calculations

- UTC(NMI) Link Uncertainty

same as document ID 112_1 (“Local clock vs. UTC”).

- Interpolation Uncertainty

  - Calculation 1

    - Standard deviation		1.7 ns

    - Comments

      - Calculated from time deviations of UTC-UTC(NMI) for every 5 days published in circular T 2010.

      - You do not have to consider this uncertainty if you calibrate equipment only on MJD 4 or 9.

  - Calculation 2

    - Frequency stability		3.9E-15 @5days

    - Comments

      - The frequency stability is calculated from UTC-UTC(NMI) for every 5 days published in circular T 2010.

      - You do not have to consider this uncertainty if you calibrate equipment only on MJD 4 or 9.

- Delay Measurement Uncertainty

same as document ID 112_1 (“Local clock vs. UTC”).

- Distribution amplifier Uncertainty

same as document ID 112_1 (“Local clock vs. UTC”).

- Cable Uncertainty

same as document ID 112_1 (“Local clock vs. UTC”).

- Counter Uncertainty

same as document ID 112_1 (“Local clock vs. UTC”).

- Counter channel differential delay

same as document ID 112_1 (“Local clock vs. UTC”).

## Uncertainty Budget table

|  | Source of Uncertainty | Uncertainty (ns) | Type | Distribution | Conversion factor | Standard Uncertainty (ns) |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | UTC(NMI) Link Uncertainty | 4.8 | B | Normal | 1 | 4.8 |
| 2 | Interpolation Uncertainty | 1.8 | A | Normal | 1 | 1.8 |
| 2 | Time Source measurement Error | 1.00 | A | Normal | 1 | 1.00 |
| 3 | Distribution amplifier Uncertainty | 0.023 | B | Rectangular |  | 0.013 |
| 4 | Cable Uncertainty | 0.02 | B | Rectangular |  | 0.01 |
| 5 | Counter Uncertainty | 0.1 | B | Rectangular |  | 0.06 |
| 6 | Counter channel differential delay | 0.5 | B | Rectangular |  | 0.3 |
| Combined Uncertainty |  |
| Expanded Uncertainty (k=2) | 11 |

## History

- 2011-12-xx first version.