# Service category 1.2.2

## Basic information

- Quantity				Time scale difference

- Instrument or Artifact			Remote clock vs. UTC

- Instrument Type or Method		GPS common view (single channel)

(Comparison against predicted UTC)

- Measurand Level or Range		-0.5 s to 0.5 s

- Measurement Conditions

  - Sampling number			720

  - Input Slew Rate			0.5 V/ns

  - Distance between 2 sites		1000 km

## Instruments

  - RF Distribution amplifier		HPDA-15RM-B (SpectraDynamics, Inc.)

  - Pulse Distribution amplifier	10188 (TimeTech GmbH)

  - Cable				FSJ1-50A (Andrew)

  - GPS receiver			NCR-37 (Japan Radio Co., Ltd.)

## Measurement configuration

## List of uncertainties

- Pseudorange Uncertainty

- Ephemeris error

- Antenna Coordinate Error

- multipath

- Source atomic clock Uncertainty

- Delay error of Ionosphere

- Delay error of Atmosphere

- GPS receiver uncertainty

- UTC(NMI) Prediction Uncertainty (30days)

- UTC(NMI) Link Uncertainty

- Time Source measurement Error

- Distribution amplifier Uncertainty

- Cable Uncertainty

## Detail of Uncertainty calculations

- Pseudorange Uncertainty

  - Receiver noise and resolution			8.00 ns

  - Sampling number					720 (total session of one day)

    - Number of GPS satellite to measure: 	8

    - Measurement number in one day: 		90

  - Comments

    - Receiver noise and resolution is calculated from one month average of standard deviation of one day data measured by a GPS receiver.

    - Number of sample is estimated from number of GPS satellite for a measurement (8) and measurement number in one day (90).

    - The multiplicative factor SQRT(2) is applied because we have to consider the uncertainty on remote and local site.

- Ephemeris error

  - Ephemeris error					+/- 2.5 m

  - Comments

    - This error is caused by a difference between the expected and actual orbital position of a GPS satellite.

    - Typical value is used as the error.

    - The multiplicative factor SQRT(2) is applied because we have to consider the uncertainty on remote and local site.

- Antenna Coordinate Error

  - Antenna Coordinate Error			+/-1 m

  - Comments

    - The multiplicative factor SQRT(2) is applied because we have to consider the uncertainty on remote and local site.

- multipath

  - multipath						+/- 1.25 m

  - Comments

    - This error is evaluated form the antenna height from ground.

    - The multiplicative factor SQRT(2) is applied because we have to consider the uncertainty on remote and local site.

    - This error will be reduced by calculation software.

- Source atomic clock Uncertainty

  - Anritsu RH401A					5.0E-15 @16 minutes

- Comments

  - This uncertainty is related to measurement for one session.

- Delay error of Ionosphere

  - Distance						1000 km

  - Difference between 2 sites			22.50 ns

  - Comments

    - Use calculated value for mid-latitude area and maximum solar activity.

    - Victor S. Zhang, “Time and Frequency Transfer Using GPS,” 30th NIST Time and Frequency Metrology Seminar, June 2005.

- Delay error of Atmosphere

  - Distance						1000 km

  - Difference between 2 sites			3.00 ns

    - Use calculated value for mid-latitude area and maximum solar activity.

    - Victor S. Zhang, “Time and Frequency Transfer Using GPS,” 30th NIST Time and Frequency Metrology Seminar, June 2005.

- GPS receiver uncertainty

NCR-37 (Japan Radio Co., Ltd.)

- Error of reciver delay				1.91 ns

- UTC(NMI) Prediction Uncertainty (30days)

Same as document ID 112_1 (“Local clock vs. UTC”).

- UTC(NMI) Link Uncertainty

Same as document ID 112_1 (“Local clock vs. UTC”).

- Delay Measurement Uncertainty

Same as document ID 112_1 (“Local clock vs. UTC”).

- Distribution amplifier Uncertainty

Same as document ID 112_1 (“Local clock vs. UTC”).

- Cable Uncertainty

Same as document ID 112_1 (“Local clock vs. UTC”).

## Uncertainty Budget table

|  | Source of Uncertainty | Uncertainty (ns) | Type | Distribution | Conversion factor | Standard Uncertainty (ns) |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Pseudorange Uncertainty | 0.42 | A | Normal | 1 | 0.42 |
| 2 | Ephemeris error | 12 | B | Rectangular |  | 7.0 |
| 3 | Antenna Coordinate Error | 4.7 | B | Rectangular |  | 2.7 |
| 4 | Multipath | 5.9 | B | Rectangular |  | 3.4 |
| 5 | Source atomic clock Uncertainty | 0 | B | Rectangular |  | 0.00 |
| 6 | Delay error of Ionosphere | 22.5 | B | Rectangular |  | 13 |
| 7 | Delay error of Atmosphere | 3 | B | Rectangular |  | 1.73 |
| 8 | GPS receiver uncertainty | 1.91 | B | Rectangular |  | 1.1 |
| 9 | UTC(NMI) Prediction Uncertainty (30days) | 14 | B | Normal | 1 | 14 |
| 10 | UTC(NMI) Link Uncertainty | 4.8 | B | Normal | 1 | 4.8 |
| 11 | Delay Measurement Uncertainty | 1.00 | A | Normal | 1 | 1.00 |
| 12 | Distribution amplifier Uncertainty | 0.023 | B | Rectangular |  | 0.013 |
| 13 | Cable Uncertainty | 0.02 | B | Rectangular |  | 0.01 |
| Combined Uncertainty |  |
| Expanded Uncertainty (k=2) | 44 |

## History

- 2011-12-xx first version.