# Service category 2.1.1

## Basic information

- Quantity				Frequency

- Instrument or Artifact			Local frequency standard

- Instrument Type or Method		Time interval measurement

- Measurand Level or Range		5MHz

- Measurement Conditions

  - Sampling number			60

  - Measurement time		1440 s

  - Amplitude				> 0.4 Vrms

- Comments

  - Total measurement time should be determined by the specification of a DUT. For example if the frequency drift of a DUT is relatively large total measurement time 1 day may be too long.

## Instruments

- RF Distribution amplifier		HPDA-15RM-B (SDI)

- cable					FSJ1-50A (Andrew)

- Time Interval Counter		SR620 (Stanford Research Systems)

## Measurement configuration

## List of uncertainties

- UTC uncertainty

- UTC-UTC(NMI) Link Uncertainty

- UTC(NMI) Uncertainty

- Interpolation uncertainty

- Distribution amplifier Uncertainty

- Cable Uncertainty

- Counter Uncertainty

## Detail of Uncertainty calculations

- UTC uncertainty

  - fractional deviation (worst value in circular T 2010)	4.6E-15

  - standard uncertainty (worst value in circular T 2010)	0.4E-15

- Comments

  - In this calculation, we assume frequency offset of UTC is always zero and fractional deviation is treated as an uncertainty.

- UTC-UTC(NMI) Link Uncertainty

  - Uncertainty uA (worst value in circular T 2010)		0.3 ns

- Comments

  - We have to consider the link uncertainty on both side of the calculation period. That is the reason why the multiplicative factor SQRT(2) is exist.

- UTC(NMI) Uncertainty

Calculate frequency difference between UTC and UTC(NMI) for every 5 days (dF) form circular T in 2010.

  - average of dF			0.3E-15

  - standard deviation of dF		5.9E-15

- Comments

  - In this calculation, we assume frequency offset of UTC(NMI) against to UTC is always zero and average of dF is treated as an uncertainty.

- Interpolation uncertainty

  - Anritsu RH401A			1.0E-15 (@ 86400s)

- Distribution amplifier Uncertainty

PD10-RM-B (SpectraDynamics, Inc.)

  - Temperature coefficient	5.0 ps/degree

  - Temperature variation	+/- 1.0 degree

  - Stability (measured)		4.0E-15 @ 10s

- Comments

  - In this calculation, we assume the temperature changes like a triangular wave and the period of the fluctuation is 1 d (= 86400 s).

  - Stability of the distribution amplifier is measured by TSC5110A (Symmetricom). The stability at 1440 seconds will be better than that of 10 seconds.

- Cable Uncertainty

FSJ1-50A (Andrew)

  - Temperature coefficient	0.036 ps/m/degree

  - Temperature variation	+/-5.0 degree

  - Cable length		60 m

- Comments

  - In this calculation, we assume the temperature changes like a triangular wave and the period of the fluctuation is 1 day (= 86400 s).

- Counter Uncertainty

SR620 (Stanford Research Systems)

- Calculation 1

  - sampling number	60

  - Counter error		100 ps

  - Measurement time	1440 s

  - Comments

    - In this calculation, we check the actual uncertainty of the time interval counter is better than the calculated value.

    - If you have to calibrate low frequency such as 1kHz, you should use strict calculation “Calculation 2”.

- Calculation 2

  - sampling number	60

  - Measurement time	1440 s

  - Timebase error	5E-13 @1s

  - short term stability	5E-13 @1s

## Uncertainty Budget table

|  | Source of Uncertainty | Uncertainty (Hz/Hz) | Type | Distribution | Conversion factor | Standard Uncertainty (Hz/Hz) |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | UTC uncertainty | 4.6E-15 | B | Normal | 1 | 4.6E-15 |
| 2 | UTC-UTC(NMI) Link Uncertainty | 1.0E-15 | A | Normal | 1 | 1.0E-15 |
| 3 | UTC(NMI) Uncertainty | 6.0E-15 | A | Normal | 1 | 6.0E-15 |
| 4 | Interpolation uncertainty | 1.0E-15 | A | Normal | 1 | 1.0E-15 |
| 5 | Distribution amplifier Uncertainty | 4.0E-15 | B | Normal | 1 | 4.0E-15 |
| 6 | Cable Uncertainty | 2.5E-16 | B | Rectangular |  | 1.5E-16 |
| 7 | Counter Uncertainty | 4.0E-13 | B | Rectangular |  | 2.3E-13 |
| Combined Uncertainty |  |
| Expanded Uncertainty (k=2) | 4.6E-13 |

## History

- 2011-12-xx first version.