# Service category 2.1.1

## Basic information

- Quantity				Frequency

- Instrument or Artifact			Local frequency standard

- Instrument Type or Method		Dual Mixer Time Difference measurement

- Measurand Level or Range		5 and 10MHz

- Measurement Conditions

  - Sampling number			86

  - Measurement time		1000 s

  - Amplitude				0.9 Vpp to 4.5 Vpp with 50 Ω

## Instruments

- RF Distribution amplifier		HPDA-15RM-B (SDI), TRAK 8451 (TRAK)

- cable					FSJ1-50A (Andrew)

- Time Interval Analyzer		TSC5110A (Symmetricom)

## Measurement configuration

## List of uncertainties

- UTC uncertainty

- UTC-UTC(NMI) Link Uncertainty

- UTC(NMI) Uncertainty

- Interpolation uncertainty

- Cable Uncertainty

- Measurement system uncertainty

## Detail of Uncertainty calculations

- UTC uncertainty

same as document ID 211_1 (“Local frequency standard”).

- UTC-UTC(NMI) Link Uncertainty

same as document ID 211_1 (“Local frequency standard”).

- UTC(NMI) Uncertainty

same as document ID 211_1 (“Local frequency standard”).

- Interpolation uncertainty

same as document ID 211_1 (“Local frequency standard”).

- Cable Uncertainty

same as document ID 211_1 (“Local frequency standard”).

- Measurement system uncertainty

  - 6E-15 @1000 s (measured)

  - The measurement system noise is measured by the setup shown in Appendix A.

  - The uncertainty usystem is the stability of the measurement result at one day time interval.

  - Distribution amplifier uncertainty is included in usystem.

## Uncertainty Budget table

|  | Source of Uncertainty | Uncertainty (Hz/Hz) | Type | Distribution | Conversion factor | Standard Uncertainty (Hz/Hz) |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | UTC uncertainty | 4.6E-15 | B | Normal | 1 | 4.6E-15 |
| 2 | UTC-UTC(NMI) Link Uncertainty | 1.0E-15 | B | Normal | 1 | 1.0E-15 |
| 3 | UTC(NMI) Uncertainty | 6.0E-15 | A | Normal | 1 | 6.0E-15 |
| 4 | Interpolation uncertainty | 1.0E-15 | A | Normal | 1 | 1.0E-15 |
| 5 | Cable Uncertainty | 2.5E-16 | B | Rectangular |  | 1.5E-16 |
| 6 | Measurement System Uncertainty | 6.5E-16 | B | Normal | 1 | 6.5E-16 |
| Combined Uncertainty |  |
| Expanded Uncertainty (k=2) | 1.6E-14 |

# Appendix A: Measurement of System Noise

- The measurement system noise was measured with setup in Figure A1.  The result is illustrated in Figure A2.

- Figure A1. Measurement setup of system noise.

- Figure A2. Frequency stability of system noise measurement.

## History

- 2011-12-xx first version.