# Service category 2.1.2

## Basic information

- Quantity				Frequency

- Instrument or Artifact			Remote frequency standard

- Instrument Type or Method		GPS common view(single channel)

- Measurand Level or Range		5 MHz

- Measurement Conditions

  - Sampling number			8

  - Averaging time			86400 s

  - Amplitude				> 0.4 Vrms

  - Distance between 2 sites		1000 km

## Instruments

- RF Distribution amplifier		HPDA-15RM-B (SDI)

- Pulse Distribution amplifier		PD10-RM-B (SpectraDynamics, Inc.)

- Cable					FSJ1-50A (Andrew)

- GPS receiver				NCR-37 (Japan Radio Co., Ltd.)

## Measurement configuration

## List of uncertainties

- Pseudorange Uncertainty

- Ephemeris error

- Antenna Coordinate Error

- multipath

- Delay error of Ionosphere

- Delay error of Atmosphere

- GPS receiver uncertainty

- UTC uncertainty

- UTC-UTC(NMI) Link Uncertainty

- UTC(NMI) Uncertainty

- Interpolation Uncertainty

- Distribution amplifier Uncertainty

- Cable Uncertainty

## Detail of Uncertainty calculations

- Pseudorange Uncertainty

  - Receiver noise and Resolution			8.0 ns

  - Sampling number					8

  - Averaging time					86400 s

  - Comments

    - Receiver noise and resolution is calculated from one month average of standard deviation of one day data measured by a GPS receiver.

    - The multiplicative factor SQRT(2) is applied because we have to consider the uncertainty on remote and local site.

    - In this measurement frequency is calculated from two time difference measurement (start and stop points), each measurement contain 8 data.

    - If DUT is a rubidium clock, it is better to use all data included in a measurement period.

- Ephemeris error

  - Ephemeris error					+/- 2.5 m

  - Averaging time					86400 s

  - Comments

    - This error is caused by a difference between the expected and actual orbital position of a GPS satellite.

    - Typical value is used as the error.

The multiplicative factor SQRT(2) is applied because we have to consider the uncertainty on remote and local site.

- Antenna Coordinate Error

  - Antenna Coordinate Error			+/- 1 m

  - Averaging time					86400 s

  - Comments

    - The multiplicative factor SQRT(2) is applied because we have to consider the uncertainty on remote and local site.

- multipath

  - multipath						+/- 1.25 m

  - Averaging time					86400 s

  - Comments

    - This error is evaluated form the antenna height from ground.

    - The multiplicative factor SQRT(2) is applied because we have to consider the uncertainty on remote and local site.

- Delay error of Ionosphere

  - Distance						1000 km

  - Difference between 2 sites			22.50 ns

  - Averaging time					86400 s

  - Comments

    - Use calculated value for mid-latitude area and maximum solar activity.

    - Victor S. Zhang, “Time and Frequency Transfer Using GPS,” 30th NIST Time and Frequency Metrology Seminar, June 2005.

- Delay error of Atmosphere

  - Distance						1000 km

  - Difference between 2 sites			3.00 ns

  - Averaging time					86400 s

  - Comments

    - Use calculated value for mid-latitude area and maximum solar activity.

    - Victor S. Zhang, “Time and Frequency Transfer Using GPS,” 30th NIST Time and Frequency Metrology Seminar, June 2005.

- GPS receiver uncertainty

NCR-37 (Japan Radio Co., Ltd.)

- Error of reciver delay				1.91 ns

- UTC uncertainty

same as document ID 211_1 (“Local frequency standard”).

- UTC-UTC(NMI) Link Uncertainty

same as document ID 211_1 (“Local frequency standard”).

- UTC(NMI) Uncertainty

same as document ID 211_1 (“Local frequency standard”).

- Interpolation uncertainty

same as document ID 211_1 (“Local frequency standard”).

- Distribution amplifier Uncertainty

same as document ID 211_1 (“Local frequency standard”).

- Cable Uncertainty

same as document ID 211_1 (“Local frequency standard”).

## Uncertainty Budget table

|  | Source of Uncertainty | Uncertainty (Hz/Hz) | Type | Distribution | Conversion factor | Standard Uncertainty (Hz/Hz) |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Pseudorange Uncertainty | 3.3E-14 | A | Normal | 1 | 4.7E-14 |
| 2 | Ephemeris error | 1.4E-13 | B | Rectangular |  | 8.1E-14 |
| 3 | Antenna Coordinate Error | 5.5E-14 | B | Rectangular |  | 3.2E-14 |
| 4 | multipath | 6.8E-14 | B | Rectangular |  | 3.9E-14 |
| 5 | Delay error of Ionosphere | 2.6E-13 | B | Rectangular |  | 1.5E-13 |
| 6 | Delay error of Atmosphere | 3.5E-14 | B | Rectangular |  | 2.0E-14 |
| 7 | GPS receiver uncertainty |  |  |  |  |  |
| 8 | UTC uncertainty | 4.6E-15 | B | Normal | 1 | 4.6E-15 |
| 9 | UTC-UTC(NMI) Link Uncertainty | 1.0E-15 | A | Normal | 1 | 1.0E-15 |
| 10 | UTC(NMI) Uncertainty | 6.0E-15 | A | Normal | 1 | 6.0E-15 |
| 11 | Interpolation uncertainty | 1.0E-15 | A | Normal | 1 | 1.0E-15 |
| 12 | Distribution amplifier Uncertainty | 4.0E-15 | B | Normal | 1 | 4.0E-15 |
| 13 | Cable Uncertainty | 2.5E-16 | B | Rectangular |  | 1.5E-16 |
| Combined Uncertainty |  |
| Expanded Uncertainty (k=2) | 3.7E-13 |

## History

- 2011-12-xx first version.