# Service category 3.4.3

## Basic information

- Quantity				Time interval

- Instrument or Artifact			Stopwatch

- Instrument Type or Method		Time interval measurement

- Measurand Level or Range		0.01 to 1E5 s

- Number of measurements         10

## Instruments

- Atomic Clock (Model : Agilent 5071A)

- Frequency Distribution Amplifier (Model : HP 5087A)

- Universal Counter (Model : Agilent 53131A)

- Stopwatch Calibrator (Model : NML-SIRIM)

## Measurement configuration

## List of uncertainties

- UTC uncertainty

- UTC-UTC(NMI) Link Uncertainty

- UTC(NMI) Uncertainty

- Cesium Uncertainty

- Cable Uncertainty

- Frequency Distribution Amplifier Uncertainty

- Universal Counter Uncertainty

- Stopwatch Calibrator Uncertainty

- Display Resolution of Stopwatch

## Details of Uncertainty Calculations

1) Using both d and u given in BIPM Circular T Section 4 (Circular T 272), we get the values of d and u given as 5.2E-15 and 0.4E-15, respectively. In this case UTC uncertainty is calculated as follows:

2)  This uncertainty is multiply by, considering that the uncertainty of TAI link both start and end points of the measurement. In this case the uncertainty is calculated as follows:

3)	From the data given in Circular T, as shown in the Appendix as Figure 1, the calculated standard deviation and average of UTC-UTC(MNFS) for MJD55214 to 55439 are 516.6E-15 and    -23.9E-15, respectively. The uncertainty is calculated as follows:

4) Values taken from Agilent 5071A (High Performance) specifications.

5) Referring to Appendix in Figure 3, we extrapolate using a line with slope τ-1 to get the uncertainty of the cable in the worst case event of ( 7(C change at 100000 s observation time.

6) From the measurement results at 100000 s observation time which are shown in Figure 5 in the appendix.

7) Calculated based on equation and value taken from Agilent 53131A specifications, as follow:

Differential Channel Error = 1.5ns

Note : The signal slew rate of the calibrator output pulse were measured using a digital oscilloscope.

Universal Counter Measurement Error = (Trigger Level Timing Error) + (Differential Channel Error) + 2(RMS Resolution) = (5.6 x 10-7) + (1.5 x 10-9) + 2(6.1 x 10-7) = ± 1.78 x 10-6s

8) The calibrator is an in-house developed unit, the uncertainty value of its trigger delay was determined from the measurement.

9) The uncertainty of the stopwatch’s time display is obtained from the smallest resolution of displayed time (typically 0.01 or 0.001 seconds).

## Uncertainty Budget Table

An example for stopwatch calibration (with 0.001 s resolution) at 100000 s measurement time and traceable to UTC.

For type A, the degree of freedom of standard uncertainty is greater or equal to 9. Whilst for type B, the degree of freedom of each standard uncertainty is believed to be large and may near to infinity.

| Uncertainty Source | Type | Value s | Distribution Factor | Sensitivity Coefficient | Standard Uncertainty s |
| --- | --- | --- | --- | --- | --- |
| UTC Uncertainty | B | 5.2E-15 x 100000 | 1 | 1 | 5.2E-10 |
| UTC-UTC(NMI) Link Uncertainty | B | 1.5E-14 x 100000 | 1 | 1 | 1.5E-9 |
| UTC(NMI) Uncertainty | B | 1.2E-12 x 100000 | 1 | 1 | 1.2E-7 |
| Cesium Uncertainty | B | 5E-13 x 100000 | 1 | 1 | 5E-8 |
| Cable Uncertainty (Worse Case : ±7 0C temperature change over 100000 s) | B | 4E-15 x 100000 | 1 | 1 | 4E-10 |
| Frequency Distribution Amplifier (± 2 0C temperature change over 100000 s) | B | 4E-15 x 100000 | 1 | 1 | 4E-10 |
| Universal Counter Uncertainty | B | 1.78E-6 |  | 1 | 1E-6 |
| Stopwatch Calibrator Uncertainty | B | 2E-7 | 1 | 1 | 2E-7 |
| Rresolution of Stopwatch | B | 5E-4 |  | 1 | 2.9E-4 |
| Combined Uncertainty = 2.9E-4 s |
| Expanded Uncertainty (k=2) = 5.8 x 10-4 s < 0.001 s (round-up) |

APPENDIX

Cable Uncertainty

This is an economical solution on creating a significant temperature changes, we used a 25 watt lighting bulb and a timer to ON and OFF the bulb. The bulb was place inside the polystyrene box with the cable under test (CUT). A thermocouple was attached to the cable’s skin and the data acquisition records its temperature at 1 s interval. The measurement set-up diagram, temperature records and measurement results (Overlapping Allan deviation) are shown as in Figure 1, 2 and 3:

Figure 1 : Measurement set-up for estimating the uncertainty of BNC cable in phase comparison in the worst case event of ( 7(C change over 1 day.

Figure 3 : Temperature measurement at 1 second interval over 1 day

Figure 4: By using extrapolation of slope τ-1, we found that the uncertainty of cable in the worst case event of ( 7(C change over 100000 s is 4 x 10-15.

Frequency Distribution Amplifier Uncertainty

Figure 2 : Temperature measurement at 1 second interval

Figure 3: By using extrapolation of slope τ-1, we found that the uncertainty of cable in worst case event of ( 7(C change over 100000 s is 4 x 10-15.

Frequency Distribution Amplifier Uncertainty

The measurement set-up diagram and measurement results (overlapping Allan deviation) are shown as in Figure 4 and Figure 5 :

Figure 4 : Measurement set-up for estimating the uncertainty of Frequency Distribution Amplifier (HP5087A) in the event of ( 2(C change over 3 days

Figure 5: The uncertainty of HP5087A in the event of ( 2(C change over 3 days

# Text boxes

start

start

stop

stop

10 MHz Reference

GPIB

Universal Counter

Frequency Distribution Amplifier

Cesium Clock

Stopwatch Calibrator

Stopwatch

Lighting Bulb

Polystyrene Box

10 MHz from Caesium Frequency Standard

Agilent 53131A

Records phase comparison measurement at

1 s gate time

External Reference In

The cable under test is a BNC-type with a 1672 cm length.

Malaysian National Frequency Standard (Agilent 5071A High Performance)

External Reference In

10 MHz and

5 MHz

1 MHz

Agilent 53131A

Records phase comparison measurement at 1 s gate time

Frequency Distribution Amplifier

(HP5087A)

1 MHz