# Service category 2.2.1

## Basic information

- Quantity				Frequency

- Instrument or Artifact			General frequency source

- Instrument Type or Method		Direct frequency measurement

- Measurand Level or Range		1Hz to 100MHz

- Measurement Conditions

  - Sampling number		173

  - Gate time			500 (s)

  - Amplitude			> 0.4 (Vrms)

## Instruments

- RF Distribution amplifier		HPDA-15RM-B (SDI)

- cable					FSJ1-50A (Andrew)

- Time Interval Counter		SR620 (Stanford Research Systems)

## Measurement configuration

## List of uncertainties

- UTC uncertainty

- UTC-UTC(NMI) Link Uncertainty

- UTC(NMI) Uncertainty

- Source atomic clock Uncertainty

- Distribution amplifier Uncertainty

- Cable Uncertainty

- Counter Uncertainty

## Detail of Uncertainty calculations

- UTC uncertainty

same as “Local frequency standard”.

- UTC-UTC(NMI) Link Uncertainty

same as “Local frequency standard”.

- UTC(NMI) Uncertainty

same as “Local frequency standard”.

- Source atomic clock Uncertainty

same as “Local frequency standard”.

- Distribution amplifier Uncertainty

same as “Local frequency standard”.

- Cable Uncertainty

same as “Local frequency standard”.

- Counter Uncertainty

SR620 (Stanford Research Systems)

  - sampling number		173

  - Gate time			500 (s)

  - frequency			1 Hz to 1E+08 (Hz)

## Uncertainty Budget table

|  | Source of Uncertainty | Uncertainty (Hz/Hz) | Type | Distribution | Conversion factor | Standard Uncertainty (Hz/Hz) |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | UTC uncertainty | 4.6E-15 | A | Normal | 1 | 4.6E-15 |
| 2 | UTC-UTC(NMI) Link Uncertainty | 1.0E-15 | A | Normal | 1 | 1.0E-15 |
| 3 | UTC(NMI) Uncertainty | 6.0E-15 | A | Normal | 1 | 6.0E-15 |
| 4 | Source atomic clock Uncertainty | 5.2E-16 | A | Normal | 1 | 5.2E-16 |
| 5 | Distribution amplifier Uncertainty | 2.3E-16 | B | Rectangular |  | 1.3E-16 |
| 6 | Cable Uncertainty | 5.0E-16 | B | Rectangular |  | 2.9E-16 |
| 7 | Counter Uncertainty | 9.0E-15 | B | Rectangular |  | 3.3E-08/f+8E-13 |
| Combined Uncertainty | 2.3E-08/f+8E-13 (f=1 Hz to 1E+08 Hz) |
| Expanded Uncertainty (k=2) | 4.6E-08/f+1.6E-12 (f=1 Hz to 1E+08 Hz) |