# Service category 2.1.1

## Basic information

- Quantity				Frequency

- Instrument or Artifact			Local frequency standard

- Instrument Type or Method		Time interval measurement

- Measurand Level or Range		5MHz

- Measurement Conditions

  - Sampling number		60

  - Measurement time	1440 (s)

  - Amplitude			> 0.4 (Vrms)

## Instruments

- RF Distribution amplifier		HPDA-15RM-B (SDI)

- cable					FSJ1-50A (Andrew)

- Time Interval Counter		SR620 (Stanford Research Systems)

## Measurement configuration

## List of uncertainties

- UTC uncertainty

- UTC-UTC(NMI) Link Uncertainty

- UTC(NMI) Uncertainty

- Source atomic clock Uncertainty

- Distribution amplifier Uncertainty

- Cable Uncertainty

- Counter Uncertainty

## Detail of Uncertainty calculations

- UTC uncertainty

  - fractional deviation (worst value in circular T 2010)	4.6E-15

  - standard uncertainty (worst value in circular T 2010)	0.4E-15

- UTC-UTC(NMI) Link Uncertainty

  - Uncertainty uA (worst value in circular T 2010)		0.3 (ns)

- UTC(NMI) Uncertainty

calculate frequency difference between UTC and UTC(NMI) for every 5 days (dF) form circular T in 2010.

  - average of dF				0.3E-15

  - standard deviation of dF		5.9E-15

- Source atomic clock Uncertainty

  - Anritsu RH401A			4.0E-15 (@ 1440s)

  - sampling number			60

- Distribution amplifier Uncertainty

PD10-RM-B (SpectraDynamics, Inc.)

  - Temperature coefficient	5.0 ps/deg

  - Temperature variation	+/- 1.0K

- Cable Uncertainty

FSJ1-50A (Andrew)

  - Temperature coefficient	0.036 ps/m/deg

  - Temperature variation	+/-5.0 K

  - Cable length		60 m

- Counter Uncertainty

SR620 (Stanford Research Systems)

  - sampling number		60

  - Counter error		100 ps

  - Measurement time	1440 (s)

## Uncertainty Budget table

|  | Source of Uncertainty | Uncertainty (Hz/Hz) | Type | Distribution | Conversion factor | Standard Uncertainty (Hz/Hz) |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | UTC uncertainty | 4.6E-15 | A | Normal | 1 | 4.6E-15 |
| 2 | UTC-UTC(NMI) Link Uncertainty | 1.0E-15 | A | Normal | 1 | 1.0E-15 |
| 3 | UTC(NMI) Uncertainty | 6.0E-15 | A | Normal | 1 | 6.0E-15 |
| 4 | Source atomic clock Uncertainty | 5.2E-16 | A | Normal | 1 | 5.2E-16 |
| 5 | Distribution amplifier Uncertainty | 2.3E-16 | B | Rectangular |  | 1.3E-16 |
| 6 | Cable Uncertainty | 5.0E-16 | B | Rectangular |  | 2.9E-16 |
| 7 | Counter Uncertainty | 9.0E-15 | B | Rectangular |  | 5.2E-15 |
| Combined Uncertainty |  |
| Expanded Uncertainty (k=2) | 2.0E-14 |