# Service category 2.1.2

## Basic information

- Quantity				Frequency

- Instrument or Artifact			Remote frequency standard

- Instrument Type or Method		GPS common view

- Measurand Level or Range		5 MHz

- Measurement Conditions

  - Sampling number		60

  - Measurement time	1440 (s)

  - Amplitude			> 0.4 (Vrms)

  - Distance between 2 sites	1000 (km)

## Instruments

- RF Distribution amplifier		HPDA-15RM-B (SDI)

- Pulse Distribution amplifier		PD10-RM-B (SpectraDynamics, Inc.)

- cable					FSJ1-50A (Andrew)

## Measurement configuration

## List of uncertainties

- Pseudorange Uncertainty

- Ephemeris error

- Antenna Coordinate Error

- multipath

- Delay error of Ionosphere

- Delay error of Atmosphere

- UTC uncertainty

- UTC-UTC(NMI) Link Uncertainty

- UTC(NMI) Uncertainty

- Source atomic clock Uncertainty

- Distribution amplifier Uncertainty

- Cable Uncertainty

## Detail of Uncertainty calculations

- Pseudorange Uncertainty

  - Receiver noise and Resolution			8.0 (ns)

  - Number of sample				16

  - Averaging time					86400 (s)

- Ephemeris error

  - Ephemeris error					5 (m)

  - Averaging time					86400 (s)

- Antenna Coordinate Error

  - Antenna Coordinate Error			1 (m)

  - Averaging time					86400 (s)

- multipath

  - multipath						2.5 (m)

  - Averaging time					86400 (s)

- Delay error of Ionosphere

  - Distance						1000 (km)

  - Difference between 2 sites			22.50 (ns)

  - Averaging time					86400 (s)

- Delay error of Atmosphere

  - Distance						1000 (km)

  - Difference between 2 sites			3.00 (ns)

  - Averaging time					86400 (s)

- UTC uncertainty

same as “Local frequency standard (Time interval measurement)”.

- UTC-UTC(NMI) Link Uncertainty

same as “Local frequency standard (Time interval measurement)”.

- UTC(NMI) Uncertainty

same as “Local frequency standard (Time interval measurement)”.

- Source atomic clock Uncertainty

same as “Local frequency standard (Time interval measurement)”.

- Distribution amplifier Uncertainty

same as “Local frequency standard (Time interval measurement)”.

- Cable Uncertainty

same as “Local frequency standard (Time interval measurement)”.

## Uncertainty Budget table

|  | Source of Uncertainty | Uncertainty (Hz/Hz) | Type | Distribution | Conversion factor | Standard Uncertainty (Hz/Hz) |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Pseudorange Uncertainty | 3.3E-14 | A | Normal | 1 | 3.3E-14 |
| 2 | Ephemeris error | 1.4E-13 | B | Rectangular |  | 8.1E-14 |
| 3 | Antenna Coordinate Error | 5.5E-14 | B | Rectangular |  | 3.2E-14 |
| 4 | multipath | 6.8E-14 | B | Rectangular |  | 3.9E-14 |
| 5 | Delay error of Ionosphere | 2.6E-13 | B | Rectangular |  | 1.5E-13 |
| 6 | Delay error of Atmosphere | 3.5E-14 | B | Rectangular |  | 2.0E-14 |
| 7 | UTC uncertainty | 4.6E-15 | A | Normal | 1 | 4.6E-15 |
| 8 | UTC-UTC(NMI) Link Uncertainty | 1.0E-15 | A | Normal | 1 | 1.0E-15 |
| 9 | UTC(NMI) Uncertainty | 6.0E-15 | A | Normal | 1 | 6.0E-15 |
| 10 | Source atomic clock Uncertainty | 5.2E-16 | A | Normal | 1 | 5.2E-16 |
| 11 | Distribution amplifier Uncertainty | 2.3E-16 | B | Rectangular |  | 1.3E-16 |
| 12 | Cable Uncertainty | 5.0E-16 | B | Rectangular |  | 2.9E-16 |
| Combined Uncertainty |  |
| Expanded Uncertainty (k=2) | 3.7E-13 |