# Service category 1.1.2

## Basic information

- Quantity				Time scale difference

- Instrument or Artifact			Local clock vs. UTC

- Instrument Type or Method		Comparison against predicted UTC

- Measurand Level or Range		-0.5 s to 0.5 s

- Measurement Conditions

  - Sampling number		100

  - Input Slew Rate		0.5 (V/ns)

## Instruments

- RF Distribution amplifier		HPDA-15RM-B (SDI)

- Pulse Distribution amplifier		PD10-RM-B (SpectraDynamics, Inc.)

- cable					FSJ1-50A (Andrew)

- Time Interval Counter		SR620 (Stanford Research Systems)

## Measurement configuration

## List of uncertainties

- UTC(NMI) Prediction Uncertainty (30days)

- UTC(NMI) Link Uncertainty

- Time Source measurement Error

- Distribution amplifier Uncertainty

- Cable Uncertainty

- Counter Uncertainty

- Counter channel differential delay

## Details of Uncertainty calculations

- UTC(NMI) Prediction Uncertainty (30days)

  - calculation 1

  - Average			7.0 (ns)

  - Standard deviation		10 (ns)

  - calculation 2

  - Maximum difference between UTC-UTC(NMI)		23.4 (ns)

  - calculation 3

  - Standard deviation of the frequency difference between UTC-UTC(NMI)	 at 30 days.	2.2E-15

- UTC(NMI) Link Uncertainty

  - uA		0.3 (ns)  (worst value in Circular T published last year)

  - uB		4.8 (ns)  (worst value in Circular T published last year)

- Time Source measurement Error

Uncertainty of delay measurement (use 5071A  Standard performance).

  - 5071A accuracy		1.0E-12

  - Measurement period	1000 (s)

- Distribution amplifier Uncertainty

PD10-RM-B (SpectraDynamics, Inc.)

  - Temperature coefficient	5.0 ps/deg

  - Temperature variation	+/-1.0 K

- Cable Uncertainty

FSJ1-50A (Andrew)

  - Temperature coefficient	0.036 ps/m/deg

  - Temperature variation	+/-5.0 K

  - Cable length		60m

- Counter Uncertainty

SR620 (Stanford Research Systems)

  - sampling number		100

  - Timebase error		5E-13 (@1s)

  - short term stability	5E-13 (@1s)

  - Input Slew Rate		0.5 (V/ns)

  - Trigger setting value	1.5 (V)

  - Time interval		1 (s)

- Counter channel differential delay

SR620 (Stanford Research Systems)

  - channel differential delay		500 (ps)

## Uncertainty Budget table

|  | Source of Uncertainty | Uncertainty (ns) | Type | Distribution | Conversion factor | Standard Uncertainty (ns) |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | UTC(NMI) Prediction Uncertainty (30days) | 14 | A | Normal | 1 | 14 |
| 2 | UTC(NMI) Link Uncertainty | 4.8 | A | Normal | 1 | 4.8 |
| 3 | Time Source measurement Error | 1.00 | A | Normal | 1 | 1.00 |
| 4 | Distribution amplifier Uncertainty | 0.01 | B | Rectangular |  | 0.01 |
| 5 | Cable Uncertainty | 0.03 | B | Rectangular |  | 0.02 |
| 6 | Counter Uncertainty | 0.1 | B | Rectangular |  | 0.06 |
| 7 | Counter channel differential delay | 0.5 | B | Rectangular | 1 | 0.5 |
| Combined Uncertainty |  |
| Expanded Uncertainty (k=2) | 30 |