# Service category 1.2.2

## Basic information

- Quantity				Time scale difference

- Instrument or Artifact			Remote clock vs. UTC

- Instrument Type or Method		GPS common view

(Comparison against predicted UTC)

- Measurand Level or Range		-0.5 s to 0.5 s

- Measurement Conditions

  - Sampling number		100

  - Input Slew Rate		0.5 (V/ns)

  - Distance between 2 sites	1000 (km)

## Instruments

  - RF Distribution amplifier		HPDA-15RM-B (SDI)

  - Pulse Distribution amplifier	PD10-RM-B (SpectraDynamics, Inc.)

  - cable				FSJ1-50A (Andrew)

## Measurement configuration

## List of uncertainties

- Pseudorange Uncertainty

- Ephemeris error

- Antenna Coordinate Error

- multipath

- Source atomic clock Uncertainty

- Delay error of Ionosphere

- Delay error of Atmosphere

- UTC(NMI) Prediction Uncertainty (30days)

- UTC(NMI) Link Uncertainty

- Time Source measurement Error

- Distribution amplifier Uncertainty

- Cable Uncertainty

## Detail of Uncertainty calculations

- Pseudorange Uncertainty

  - Receiver noise and resolution			8.00 (ns)

  - Number of sample				720

- Ephemeris error

  - Ephemeris error					5 (m)

- Antenna Coordinate Error

  - Antenna Coordinate Error			1 (m)

- multipath

  - multipath						2.5 (m)

- Source atomic clock Uncertainty

  - Anritsu RH401A					5.0E-15 (@ 16 min)

- Delay error of Ionosphere

  - Distance						1000 (km)

  - Difference between 2 sites			22.50 (ns)

- Delay error of Atmosphere

  - Distance						1000 (km)

  - Difference between 2 sites			3.00 (ns)

- UTC(NMI) Prediction Uncertainty (30days)

Same as “Local clock vs. UTC”.

- UTC(NMI) Link Uncertainty

Same as “Local clock vs. UTC”.

- Time Source measurement Error

Same as “Local clock vs. UTC”.

- Distribution amplifier Uncertainty

Same as “Local clock vs. UTC”.

- Cable Uncertainty

Same as “Local clock vs. UTC”.

## Uncertainty Budget table

|  | Source of Uncertainty | Uncertainty (ns) | Type | Distribution | Conversion factor | Standard Uncertainty (ns) |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Pseudorange Uncertainty | 0.42 | A | Normal | 1 | 0.42 |
| 2 | Ephemeris error | 12 | B | Rectangular |  | 7.0 |
| 3 | Antenna Coordinate Error | 4.7 | B | Rectangular |  | 2.7 |
| 4 | Multipath | 5.9 | B | Rectangular |  | 3.4 |
| 5 | Source atomic clock Uncertainty | 0 | B | Rectangular |  | 0.00 |
| 6 | Delay error of Ionosphere | 22.5 | B | Rectangular |  | 13 |
| 7 | Delay error of Atmosphere | 3 | B | Rectangular |  | 1.73 |
| 8 | UTC(NMI) Prediction Uncertainty (30days) | 14 | A | Normal | 1 | 14 |
| 9 | UTC(NMI) Link Uncertainty | 4.8 | A | Normal | 1 | 4.8 |
| 10 | Time Source measurement Error | 1.00 | A | Normal | 1 | 1.00 |
| 11 | Distribution amplifier Uncertainty | 0.01 | B | Rectangular |  | 0.01 |
| 12 | Cable Uncertainty | 0.03 | B | Rectangular |  | 0.02 |
| Combined Uncertainty |  |
| Expanded Uncertainty (k=2) | 44 |